Physical Characterization of Solid-Liquid Slurries at High Weight Fractions Using Optical and Ultrasonic Methods
DE-FG02-05ER63990 

UW BN 61-1557 

Final Report 
Lead Principal Investigators:

 L.W. Burgess and A.M. Brodsky

Center for Process Analytical Chemistry

University of Washington

Seattle, Washington

(206) 543-0579

Lloyd@cpac.washington.edu
Co-Investigators:

P.D. Panetta, B. Tucker, and K. Denslow
Pacific Northwest National Laboratory

Richland, Washington

(509) 372-6107

Paul.Panetta@pnl.gov
Graduate Student:

S.L. Randall

Center for Process Analytical Chemistry

University of Washington

Seattle, Washington

(206) 616-1492

sr33@u.washington.edu
Executive Summary
Remediation of highly radioactive waste is a major technical and programmatic challenge for the DOE.  Rapid, on-line physical characterization of highly concentrated slurries is required for the safe and efficient remediation of 90 million gallons of high level radioactive waste (HLW), sodium bearing waste, and mixed waste.  The research presented here, describes a collaborative effort between Pacific Northwest National Laboratory (PNNL) and the University of Washington to directly address the need for rapid on-line characterization of the physical properties of HLW slurries during all phases of the remediation process, from in-tank characterization of sediments to monitoring of the concentration, particle size, and degree of agglomeration and gelation of slurries during transport. Near-surface characterization of the slurry flow in the particle size range from nanometer to micrometer is examined using optical low coherence reflectometry.  Volumetric characterization at depths in the slurry flow, up to several centimeters in the particle size range from the micrometer to millimeter, is realized by utilizing ultrasonic backscatter and diffuses fields. One of the strengths, the teaming up of significant talents in both experimental and theoretical optics (University of Washington) and in ultrasonics [Pacific Northwest National Laboratory (PNNL)] provides a synergistic approach to integrate these complimentary techniques.  One of the benefits of this combined approach is the physical characterization of HLW over a concentration and particle size range that is broader than can be achieved with today’s technology. This will avoid a costly increase in waste stream volume due to excess dilution, and will lessen chance of plugging pipes that could shut down expensive processing lines.
Research Progress and Implications
The project, initially funded under DE-FG07-01ER63267, received continuing project support under DE-FG02-05ER63990.  This report represents progress during the initial and renewal funding periods, and a subsequent no cost extension which aided us in bridging a gap (approximately six months) between the initial funding period and the beginning of the renewal period. Slurry systems with mean particle sizes from ~0.01um to 800 um have been studied at concentration ranging from very dilute (< 1 wt %) to 40 wt %.  The interrelationships between the attenuation, the direct backscattered field, and the diffusely propagating field have been investigated.  These fields are rich with information and unraveling their interrelationships provides the robust methodology needed to characterize these HLW slurries and enhance the clean up process at the DOE facilities.  

This integration of both optic and ultrasonic techniques has been unique and exciting because it offers a combined methodology that exploits the best features of these two powerful techniques.  The teaming of significant expertise in both experimental and theoretical optics (University of Washington) and in experimental and theoretical ultrasonics (PNNL) has provided a synergistic approach to coupling these complementary techniques. This combination of capabilities has allowed characterization to be performed over a broader range of particle sizes and concentrations than can be accomplished with existing techniques.  
There have been three tasks:  1) develop new optical and acoustic scattering measurements to provide the fundamental science needed for successful device development and implementation, 2) develop theories that describe the interrelationship between wave propagation and the physical properties of the slurry, and 3) perform inversions of the theories and compare them with the experimental measurements to non-intrusively characterize slurries.
Optical Measurements
The optical measurements are based on the investigation of slurries with Optical Low-Coherence Reflectometry (OLCR).  Optical Low Coherence Reflectometry (OLCR) is a non-destructive optical technique that can be used to measure 180o coherent backscatter of light incident on a heterogeneous medium. OLCR is a white-light interference technique originally developed for applications in the optoelectronics test and measurement arena.  The instrumental design is an autocorrelator to detect interference fringe position and intensity, using a broadband LED source [3].  The detection of only coherent light allows us to extract characteristic information about the sample, such as particle size and concentration.  

Tailing decay profiles, in the form of reflectance vs. photon dwell distance (time spent within the matrix), are generated with OLCR from highly scattering samples as shown in Figure 1a.  As described in Equation 1, reflectance is related to the proportion of incident light which is returned in phase to the optical fiber Changes in particle size and concentration may be extracted from the OLCR decay profiles of highly concentrated slurries and emulsions, such as high-level waste. We have shown OLCR to be a valuable tool for monitoring changes in particle size over several orders of magnitude at high concentration.

OLCR is a white-light interferometric technique that incorporates a broadband light source with a classical Michelson interferometer.  In previous works1-3 we have demonstrated that OLCR is an effective tool for nondestructive analysis of multiscattering systems. During our work we have studied many aspects of the OLCR system for slurry property monitoring. Initial work was performed using a Hewlett-Packard OLCR instrument and evolved to use the more flexible Optiphase instrument. The probe design and development for the Optiphase OLCR instrument has been a non-trivial aspect of the instrumentation. The instrument design requires a triggering reflection, which we have chosen to be from the fiber end face without any contact from the sample.  By adding the non-contact limitation, the fiber end face reflection can be used not only to trigger the instrumentation, but also to serve as a reference to correct for source fluctuations and allows us to track the matrix refractive index. For solid-liquid slurries, the probe has been designed to consist of standard single mode optical fiber, followed by a small air gap, and then a glass window angled at 8( relative to the fiber end face.  This design results in triggering and reference reflections from the interface between the fiber end face and air, and the interface between air and the inside surface of the glass window.  The reflection resulting from the outside surface of the glass window, which is in contact with the sample, will vary in amplitude based upon the refractive index difference between the window and the sample. The probe design is illustrated as Figure 1 and the resulting OLCR profiles from the solvent series are shown in Figure 1B. The fourth reflection is a resonance effect which occurs due to the length of the air cavity within the probe.  In later iterations, the air cavity length was decreased and the resonance reflection disappeared. The amplitude ratio of the reflection from the air-inside window interface to the reflection of the outside window-sample interface can be related to the refractive index of the sample. In the Rayleigh scattering regime, the refractive index can be directly related to the concentration of dispersed material of known refractive index, and can be diagnostic during agglomeration as particles transition from one scattering regime to the other. This feature aids us in deconvoluting dynamic changes in the slurry system.

[image: image1.jpg]Coherent Reflectance (4B)

8§08 & 8 2 8 % 4 &

H

Phetun Dwell Distance (mm)




Figure 1. A) Simplified probe design consisting of bare single-mode fiber, the air cavity, and the glass window at an 8 degree angle. B) The resulting four reflections from a series of solvents with varying refractive indices.

In the optical experimental work we have used standard monodispersed polystyrene nanospheres purchased from Duke Scientific Corporation to act as a model system for theory and experimental development over a large dynamic range of particle sizes.  The shape and intensity of the OLCR profiles changes as the measured particle size transitions from the Rayleigh regime to the Fraunhofer regime, as shown in Figure 2. The spike at photon dwell distances around 3 mm are due to a source artifact and are not related to the sample.
[image: image14.emf]10 MHz transdcuer

0.36

0.38

0.40

0.42

0.44

0.46

0.48

0.50

0.52

8 8.5 9 9.5 10

Frequency (MHz)

Attenuation (nepers/cm)

Water Run2 Run3 1% 2% 3% 4% 5%

0% 1um Contamination: 2 Runs

1% - 5% vol% of 5wt% 1um

(thus 99% - 95% vol% of 5wt% 0.05um)

1.5 PPM

6.8 PPM

 
Figure 2.  OLCR profiles for polystyrene sphere standards for diameters 20 nm to 1745 nm.
 One of the most interesting results of this work is the observation of so-called Mie resonances when the particle radii of a randomly distributed system are comparable with the wavelength of light.  This non-monotonic oscillation of backscattered intensity is dependent upon optical parameters, such as dielectric contrast between the medium and the particles.
Advances in the optical theory of multiple scattering have been made, specifically to begin to extract information pertaining to system properties, such as particle size, concentration, and polydispersity, from all regions of the OLCR profile.  Theoretical estimations of the experimental amplitude oscillation are described in the following equation:
S – S0 = Cρ0 [0.25 + 1/y2 + 2(1-cos(y))/y4 – sin(y)/y + (1-cos(y))/y2 – 2siny/y3]
→ Cρ0 (R/λ)2 (4πΔnM2/3)2

                                for RΔ/λ<<1 (Rayleigh Limit)


(1)
where S is the total signal, S0 is a combination of an adjustment for the baseline signal and an indication of the non-uniformity of the solution, C is an instrumentally-based constant, ρ0 is the volume fraction of particles in the studied media, y = 4πRΔεM/λ with R as the particle radius, Δ as the optical contrast between the particles and media, εM is the dielectric function of the media, and λ as the wavelength of light in a vacuum.  The described results are interesting as a step toward advancement in the understanding of the theoretically and practically important problem of wave propagation in multiscattering random media.  


Analysis of two polystyrene suspension series with the same range of particle sizes but different volume fractions led to a few interesting comparisons.  The calculated effective refractive indices of both suspensions were equal, 1.3585 RI units, based upon the calibration established by the peak ratios from the solvent series described above.  The theoretical expression as described in the sections above suggested that an increase by a factor of ten in volume fraction led to an increase by a factor of approximately ten for the amplitude of Mie resonances, and this was confirmed by comparison of experimental data.  The two experimental Mie resonance curves from the 1% (v/v) and 10% (v/v) suspension series are shown as Figure 3. Manuscripts has been published to describe both the experimental and theoretical Mie resonance results4,5.  
[image: image2.wmf]0.00E+00

2.00E-06

4.00E-06

6.00E-06

8.00E-06

1.00E-05

0

500

1000

1500

2000

Particle Diameter (nm)

S - S

0

10% (v/v)

1% (v/v)

1.0 x 10

-

5

8.0 x 10

-

6

4.0 x 10

-

6

2.0 x 10

-

6

0

S 

–

S

0

6.0 x 10

-

6


Figure 3. Comparison of experimental Mie resonance curves from a 1% (v/v) polystyrene suspension series (circles) and a 10% (v/v) polystyrene suspension series.
Both standard polystyrene spheres and DOE waste simulants have been used for polydispersed measurements.  We have reported on OLCR measurement capabilities in a more complex and realistic matrices including bimodal mixtures of individual polydispersed alumina suspensions. These were easily distinguishable from each other when measured individually at 30 wt %, and the trend of changing average particle size can be easily followed through the profile amplitude.  This result is very exciting and extremely encouraging towards our ability to track changing dimensions in a complex, polydispersed, multi-modal surrogate slurry.  The alumina component serves as an excellent surrogate waste model for both static and dynamic studies. In Figure 4 below, reproducible scattering profiles are shown for S-23 alumina in a range of concentrations from a 1 to 40 weight percent.
[image: image3.wmf]S-23 Alumina Slurry - Each trace represents the average of 100 sequential scans (~30 sec)       

5 traces at each concentration.

-82

-77

-72

-67

-62

-57

-52

0.3

0.5

0.7

0.9

1.1

1.3

1.5

1.7

1.9

2.1

2.3

Photon Dwell distance (mm)

Reflectance (dB)

1% 1

1% 2

1% 3

1% 4

1% 5

10% 1

10% 2

10% 3

10% 4

10% 5

30% 1

30% 2

30% 3

30% 4

30% 5

40% 1

40% 2

40% 3

40% 4

40% 5

 
Figure 4. SpaceRite S-23 Alumina (Gibbsite) which has a broadly distributed 7.5 micron mean particle size

We have developed a greater understanding of how the complete surrogate slurry will act through the analysis of bimodal monodispersed polystyrene mixtures.  Bimodal polystyrene solutions were prepared by mixing monodispersed 1% w/w standards in varying amounts to form the desired average size or size ratios as shown in the table in Figure 4. The five bimodal samples were analyzed which had a mean particle size of 60 nm, but different compositions.  The experimental method used was consistent with that from the standard monodispersed systems. The decay curves of the resulting OLCR profiles are very similar in intensity and slope and clear relationships between the contributions from the bimodal components and the regions of the profiles cannot be easily determined. 

Differences between the bimodal system composition arise however, when the fluctuations of an individual signal from the mean are analyzed throughout the entire OLCR profile.  These fluctuations are not noise-based and do not average out with multiple scans in the same location of a solid sample.  Rather, they are due to fixed heterogeneities within the system and contain information regarding particle clustering.  The C-value for fluctuation analysis is calculated by 

C = <J(L)J(L+L)> / <J(L+L)><J(L)> -  1


(2)
where J(L) is the OLCR signal at a wavepacket dwell distance of L and L is indicative of the step size or correlation distance in mm.  Log C2 is plotted for each L, or correlation distance in mm. Figure 5 illustrates that with increased correlation distance, the distinction between the samples of the various compositions becomes more pronounced.  First, variation in the signal is measured over each scan from the mean over the decay length of the profile.  Then, by ‘stepping through’ the OLCR profiles and correlating the intensity of individual data points to their neighbors’, a sense of a correlation distance, L, is found, which is related to the composition of the matrices. Fluctuations are especially important for monitoring phase transitions, such as gelation.  The OLCR profile of a liquid solution is relatively smooth, due to Brownian motion causing an averaging of the signal fluctuations.  However, as a sample transitions into a more solid state, Brownian motion slows and the fluctuations in the OLCR profile increase.  
[image: image4.jpg](A,.',':, s:,',',.’:’ 100nm| 80nm | 60nm | 40nm 20 nm
100 A 100%

8 | B 100%

60 | 4 100%

| D 100%

20| E 100%
80 | F 50% 50%

8| G 66.7% 333%

80 | H 75% 25%
60 | 5 50% 50%

60 | 1 50% 50%
60 | 3 66.7% 333%
6 | 2 333% 66.7%

0| M 50% 50%
| N 333% 66.7%
| o 25% 75%





[image: image5.wmf]Correlation Distance, 

Δ

L 

(mm)

-

9

-

8

-

7

-

6

-

5

-

4

0

0.005

0.01

0.015

0.02

Average Log (C

2

)

1

2

3

4

5

Correlation Distance, 

Δ

L 

(mm)

-

9

-

8

-

7

-

6

-

5

-

4

0

0.005

0.01

0.015

0.02

Average Log (C

2

)

1

2

3

4

5



Figure 5.  The experimental design and a figure illustrating how the differentiation between bimodal suspensions (indicated in red) increases with increasing correlation distance.

However, as a sample transitions into a more solid state, Brownian motion slows and the fluctuations in the OLCR profile increase.  This is because only limited numbers of random particle distributions are available for analysis by the wavepacket.  When the system under test becomes solid, only one distribution is available and the fluctuations will become very stable. 
Gelation, agglomeration and sedimentation have all been shown to occur by adjusting parameters of the alumina system (i.e. pH, flow rate, temperature). Agglomeration is easily recognized as an increase in particle size. The dynamics of the sedimentation can easily be followed by continuously monitoring both the amplitude and the dwell distance of the scatter profile.  This can easily be demonstrated as shown in Figure 6.

[image: image6.emf]WYS050202 - OLCR of Settling TiO2 (4 hours)

-75

-70

-65

-60

-55

-50

0.5 1 1.5 2 2.5 3 3.5 4 4.5

Wave Packet Dwell Distance/mm

dB


Figure 6 A sequence of scattering profiles collected over a four-hour period observing a fine titanium dioxide slurry.  (much slower sedimentation rate relative to the alumina surrogates). Gelation is a much more interesting situation. Two of most widely observed alumina phases in the tank waste are gibbsite and boehmite, represented by HiQ-10 (boehmite) and the S-23 (gibbsite) surrogates.  The boehmite phase will undergo a crystal into gel phase transition upon acidification.  In the gel phase, the system takes on a long-range order which all but eliminates 180 degree coherent backscatter.  This takes place rather rapidly upon addition of small aliquots of concentrated acid (HCl) to a stirred slurry matrix.  The observed backscatter profile changes had approximately 20 second intervals is shown in Figure 7. 
[image: image7.emf]Gellation Dynamics -100 microliter additions of 12M HCL to 25mL 

15 wt% HiQ-10 +approximate seconds after each addition

-80

-79

-78

-77

-76

-75

-74

-73

-72

-71

-70

0.4 0.6 0.8 1 1.2

Photon Dwell Distance (mm)

Amplitude (dB)

15 wt% HiQ-10

1st addition+5sec

1st addition+25

1st addition+110

2nd addition +10

2nd add +50

2nd addition +110

3rd addition +10

3rd addition +30

3rd addition +50

3rd addition +110

4th addition +110

[image: image8.jpg]



Figure 7.  The backscatter decay profiles resulting as we follow the gelation dynamics of HiQ-10 and the resultant gel after 400 microliters of 12M HCl.
Overall, we have made significant progress in the analysis of signals obtained from multiple scattering events.  We have observed the appearance of Mie resonances, a non-monotonic oscillation of backscattered intensity dependent on optical parameters, over the interval where radii are comparable to the wavelength. The results above clearly indicate that it is possible to extract important static and dynamic parametric information from dense scattering systems with OLCR.  Between the refractive index information, fluctuation distributions, amplitude variation, and decay rate, we are able to extract many of the characteristics of particle systems.  If all the variables are changing simultaneously further resolution of the parameter convolutions can be addressed by utilizing the synergy of the project and incorporating information from the ultrasound data.  

Ultrasonic Measurements
During this period PNNL has been developing ultrasonic measurement methodologies to characterize solid liquid suspensions at high concentrations.  A schematic of the measurement system is shown in Figure 8a with representative signal used for calculating the attenuation and speed of sound in Figure 8b, the backscattering in Figure 8c, and the diffuse field in Figure 8d.  The measurement cell is composed of a ring with opposing transducers to send and receive the ultrasonic fields.  The speed of sound is simply a measure of the transit time across the chamber and the attenuation is a measure of the amount of energy lost due to scattering and absorption as the sound traverses the chamber.  The backscattering is a measure of the amount of energy that is returned to the transmitting transducer.  The diffuse field develops after the ultrasonic field has undergone many scattering events with the particles in the fluid and the walls of the chamber.   For the attenuation, backscattering and the diffuse field, analysis is performed to produce an ultrasonic response as a function of frequency.  The effects from the measurement system and the transducer are removed from the received signal by utilizing an appropriate reference signal.  Details can be found in publications that resulted from this work. 
[image: image9.emf]-2

-1.5

-1

-0.5

0

0.5

1

1.5

2

0 20 40 60 80 100 120 140

Time  (us)

Backscattered Amplitude  (Volts)

(b) Attenuation and Velocity

(d) Diffuse field

-0.5

0.0

0.5

0 1000 2000

Time (us)

Amplitude (volts)

-2

-1

0

1

2

67.5 68 68.5 69

Time (us)

Amplitude (Volts)

0 cells/mL

1.9 X 10

8

 cells/mL

Low concentration High concentration

(c) Backscattering (70 um particles)

40 wt%

5 wt%

15 wt%

(a)

Attenuation, Velocity

Transmitter/

Receiver 

Receiver for diffuse 

field and off axis 

backscattering

Attenuation, Velocity

Transmitter/

Receiver 

Receiver for diffuse 

field and off axis 

backscattering

-2

-1.5

-1

-0.5

0

0.5

1

1.5

2

0 20 40 60 80 100 120 140

Time  (us)

Backscattered Amplitude  (Volts)

(b) Attenuation and Velocity

(d) Diffuse field

-0.5

0.0

0.5

0 1000 2000

Time (us)

Amplitude (volts)

(d) Diffuse field

-0.5

0.0

0.5

0 1000 2000

Time (us)

Amplitude (volts)

-2

-1

0

1

2

67.5 68 68.5 69

Time (us)

Amplitude (Volts)

0 cells/mL

1.9 X 10

8

 cells/mL

Low concentration High concentration

-2

-1

0

1

2

67.5 68 68.5 69

Time (us)

Amplitude (Volts)

0 cells/mL

1.9 X 10

8

 cells/mL

Low concentration High concentration

(c) Backscattering (70 um particles)

40 wt%

5 wt%

15 wt%

(a)

Attenuation, Velocity

Transmitter/

Receiver 

Receiver for diffuse 

field and off axis 

backscattering

Attenuation, Velocity

Transmitter/

Receiver 

Receiver for diffuse 

field and off axis 

backscattering


Figure 8.  (a) Schematic of ultrasonic measurement apparatus, (b) through transmitted RF signal for attenuation and velocity determination, (c) directly backscattered signal, and (d) diffuse field signal.
To characterize polydispersed specimens, the ultrasonic attenuation as a function of frequency was measured for the Al2O3 specimens shown in Table 1.  The attenuation as a function of frequency is shown in Figure 9.  In addition to the attenuation increasing as the particles were added, the dependence on frequency also changed.  The change in frequency dependence can be explained by considering the frequency dependence of the attenuating mechanisms on the product of frequency and particle size as shown in the right side of Figure 9.  Specifically, as the average particle size increased, the frequency response increased following the curvature between the frequency to the ½ power and frequency to the 4th power.  The precise measurements show sensitivity to small changes in concentration and average particle size.  

To further characterize polydispersed specimens, a second set of specimens were prepared as described in Table II, where the concentration was held constant wile the particle size distribution was altered.  Precise ultrasonic attenuation measurements were performed as a function of frequency to determine the sensitivity to small changes in the particle size distribution.  It is significant that the PNNL measurements of the ultrasonic attenuation were made with such high precisions that we were able to measure changes as small as a 1.5 PPM addition of 1 um particles to a slurry of 0.05 um particles.  Further measurements utilizing the backscattering and the diffuse field are being performed and will be reported in subsequent reports.
[image: image10.emf]-85

-80

-75

-70

-65

-60

-55

-50

0.00 0.10 0.20 0.30 0.40

Photon Dwell Distance (mm)

Coherent Reflectance (dB)

Air 

Cavity

Glass 

Thickness

Double 

Reflection

B.

Fiber

Air

Glass

Sample

…

A.

Fiber

Air

Glass

Sample

…

A.

Table I.  Polydispersed specimens where both concentration and the particle size distribution were altered.
[image: image11.emf]1091.5

2.18E+11 33wt% 1um particles

766.3

1.53E+11 25wt% 1um particles

270.1

5.39E+10 10wt% 1um particles

129.8

2.59E+10 5wt% 1um particles

103.2

2.06E+10 4wt% 1um particles

76.7

1.53E+10 3wt% 1um particles

50.6

1.01E+10 2wt% 1um particles

25.2

5.02E+09 1wt% 1um particles

PPM of 1 um 

particles

Number of 1 um 

particles/mL of 

suspension

Began with 900 mL of 0.05 

um suspension

MIXTURES OF 1 um DIAMETER PARTICLES IN A 5 WT% 0.05 um DIAMETER 

PARTICLE BASE SUSPENSION

1091.5

2.18E+11 33wt% 1um particles

766.3

1.53E+11 25wt% 1um particles

270.1

5.39E+10 10wt% 1um particles

129.8

2.59E+10 5wt% 1um particles

103.2

2.06E+10 4wt% 1um particles

76.7

1.53E+10 3wt% 1um particles

50.6

1.01E+10 2wt% 1um particles

25.2

5.02E+09 1wt% 1um particles

PPM of 1 um 

particles

Number of 1 um 

particles/mL of 

suspension

Began with 900 mL of 0.05 

um suspension

MIXTURES OF 1 um DIAMETER PARTICLES IN A 5 WT% 0.05 um DIAMETER 

PARTICLE BASE SUSPENSION


Figure 9.  Attenuation as a function of frequency for polydispersed alumina particles in water where both the concentration and the particle size distribution were altered.

Table II  Polydispersed specimens where the particle size distribution was altered while holding the concentration constant.

[image: image12.emf]0

0.2

0.4

0.6

0.8

1

1.2

0 1 2 3 4 5 6 7 8 9

Frequency (MHz)

Attenuation (nepers/cm)

5%_0.05um only

50 PPM

130 PPM

103 PPM

77 PPM

25 PPM

1100 PPM

766 PPM

270 PPM

0

1

2

3

0 0.5 1 1.5

frequency

Attenuation

Viscous

Multiple

Scattering

Inertial

ka

f

4

f

1/2

f

2

Frequency * particle size        


As part of this project, we are advancing this understanding to better utilize the backscattering and diffuse field measurements to characterize solid liquid suspensions.  The backscattering and diffuse field measurements are especially appealing because of the relative simplicity of the measurement and the theoretical description of the scattering processes.  Advancements to the measurements performed thus far are being further analyzed to increase our understanding of the mechanisms which govern the ultrasonic scattering and absorption contributions to these measurements.  The results are currently being prepared for a patent application.  These studies provide new advancements in the understanding and applications of the ultrasonic backscattering to characterize high concentration solid-liquid suspension.
We were able to attempt one set of experiments to analyze a complex matrix composed of multiple surrogate components with both optical and acoustic techniques. The optical probe consisted of a 1/4 inch stainless steel tube terminated as shown in Figure 1a. Because of the design of the experimental acoustic apparatus (Figure 8a) access to the stirred matrix was carried out by inserting the optical probe into the center of the ring being careful not to obstruct the acoustic pathway.  Unfortunately, the vortex created by this quarter inch tube, even with very moderate stirring rates, caused sufficient gas to the entrained into the matrix that the acoustic backscatter signals disappeared!  This entrained gas had no effect on the optical signals.  This can be readily understood by the fact that the acoustic probes sample from a much larger volume of the matrix than the optical probe. Reengineering the apparatus was not feasible in the remaining project period.  As a result, we were unsuccessful in obtaining dual sensor simultaneous responses.

[image: image13.emf]6.8

6.77E-06 1.89E+14

2.00E+14

1.48E+05 1.28E+09 5

5.4

5.39E-06 1.91E+14

2.00E+14

1.85E+05 1.03E+09 4

3.9

3.90E-06 1.93E+14

2.00E+14

2.57E+05 7.52E+08 3

2.7

2.70E-06 1.95E+14

2.00E+14

3.70E+05 5.27E+08 2

1.5

1.51E-06 1.97E+14

2.00E+14

6.61E+05 2.98E+08 1

PPM

# 1um/ # 50 nm

Number of 

small+large

particles/m

L now:

Number of 50 nm 

particles/m

L of 

suspension

Number of 50 nm 

particles to 

every one 1 

um 

particle:

Number of 1um 

particles/m

L of 

suspension

:

vol% of 1um 

5wt%

6.8

6.77E-06 1.89E+14

2.00E+14

1.48E+05 1.28E+09 5

5.4

5.39E-06 1.91E+14

2.00E+14

1.85E+05 1.03E+09 4

3.9

3.90E-06 1.93E+14

2.00E+14

2.57E+05 7.52E+08 3

2.7

2.70E-06 1.95E+14

2.00E+14

3.70E+05 5.27E+08 2

1.5

1.51E-06 1.97E+14

2.00E+14

6.61E+05 2.98E+08 1

PPM

# 1um/ # 50 nm

Number of 

small+large

particles/m

L now:

Number of 50 nm 

particles/m

L of 

suspension

Number of 50 nm 

particles to 

every one 1 

um 

particle:

Number of 1um 

particles/m

L of 

suspension

:

vol% of 1um 

5wt%


Figure 10.  Attenuation as a function of frequency for polydispersed alumina particles in water, where only the particle size distribution was altered.
Project research has resulted in numerous presentations at the International Particle Size Analysis 2003 Conference (Harrogate, UK), at the Center for Process Analytical Chemistry Semi-Annual Meetings, at the Review of Progress in Quantitative Nondestructive Evaluation 2002 and 2003, at the International Forum for Process Analytical Chemistry 2004, at the IEEE Instrumentation and Measurement Technology Conference 2003, at the National American Chemical Society Conference in Anaheim 2004, and the Review of Progress on Quantitative NDE, 2004.  Publications as a result of this research include Brodsky and Burgess [2003], Randall et al. [2003], Randall et al. [2005] , Brodsky et. al.[2007] Panetta et al. [2003], Panetta et al. [2004] , and the Ph.D. Dissertation of Summer Lockerbie Randall [2004].
References
1. Thurber, S.R., A.M. Brodsky, and L.W. Burgess. “Characterization of Random Media by Low Coherence Interferometry.”  Applied Spectroscopy, 54(10), 1506-1514 (2000).

2. Brodsky, A., Thurber, S.R., and L.W. Burgess.  “Low-Coherence Interferometry in Random Media I.  Theory.”  Journal of the Optical Society of America A, 17(11), 2024-2033 (2000).Thurber, S.R.,    

3. Burgess, L.W., A. Brodsky, and P.H. Shelley.  “Low-Coherence Interferometry in Random Media II. Experiment.”  Journal of the Optical Society of America A, 17(11), 2034-2039  (2000).
4. S. L. Randall, A. M. Brodsky, and L. W. Burgess “Manifestation of Mie Resonances in Light Scattering from Multiscattering Media”, Inter. J. of Mod. Phys B, V. 19, No. 4  (2005)
5. Anatol M. Brodsky, Gordon T. Mitchell, Summer L. Ziegler, and Lloyd W. Burgess, “Coherence loss in light backscattering by random media with nanoscale non uniformities”, Phys. Rev. E 75, 046605 (2007)
6. Randall, S.L., A.M. Brodsky, L.W. Burgess, and R.L. Green.  “Optical Low-Coherence Reflectometry for Nondestructive Process Measurements,” 29th Annual Review of Progress in Quantitative Nondestructive Evaluation (QNDE), Bellingham, WA (2003).

7. Brodsky, A. and L. Burgess.  .  “Theoretical Study of the Coherent Backscattering of Light by Disordered Media.” Intl. J. Modern Physics B, 17(3), 337 (2003).
8. S. L. Randall, “Development and Utilization of Optical Low Coherence Reflectometry for the Study of Multiple Scattering in Randomly Distributed Solid-Liquid Suspensions”, Ph.D. Dissertation, University of Washington, December, (2004)

9. Golcar, G.R.,  K.P. Brooks, J.G. Darab, J.M. Davis, and L.K. Jagoda. 2000. Development of Inactive High Level Waste  Envelope D Simulants for Scaled Crossflow Filtration Testing, PNWD-3042.

10. Panetta, P.D., B.J. Tucker, R.A. Pappas, and S. Ahmed,” Characterization of solid liquid suspension utilizing ultrasonic measurements”, 29th Annual Review of Progress in Quantitative Nondestructive Evaluation -2003
11. Panetta, P.D., B. Tucker R. A. Pappas, and S. Ahmed, “Characterization of Solid Liquid Suspensions Utilizing Ultrasonic Measurements” IEEE Instrumentation and Measurement Technology Conference (in press)

PAGE  
3

